pH Meter (F20) User Manual
Basic Information
Band/Model	METTLER TOLEDO / F20
No. of property	3100801002-1284, 3100801002-1285

Performing a 3-point calibration ( pH 4,7,10)
1. Press and realease to switch on the instrument.
2. Place the electrode in a calibration buffer.
3. Press . During measurement the pH value based on the previous calibration is shown. The instrument stops measuring when the signal is stable. (auto endpoint)
4. Proceed with next calibration point and go to performing the next point calibration.
5. Rinse the electrode with deionize water.
6. Place the electrode in the next calibration buffer and press . The instrument stops measuring when the signal is stable. (auto endpoint)
7. Proceed with next calibration point and go to performing the next point calibration.
8. Rinse the electrode with deionize water.
9. Place the electrode in the next calibration buffer and press . The instrument stops measuring when the signal is stable. (auto endpoint)
10. Press to finish and save the 3-point calibration.
11. Both slope and offset are shown on the right side of the display.
Performing a measurement
1. Make sure that the pH reading mode is selected.
2. Place the electrode in the sample and press to start the measurement.
3. The decimal point blinks, If the automatic endpoint is selected, and the signal has stabilized, the display freezes, appears and the decimal point stops blinking.
Press for 3seconds and release to switch off the instrument.

Notice
1. Make sure pH electrodes are clean and always keep filled with 3M KCL.
2. If the electrode slope falls rapidly, soak the tip of the electrode overnight in 0.1M HCl.